Luna GUI / Luna Instrument Controller Command Interface Definition
Command Format:

ID: 10 bytes (0-9)
Right justified, zero filled, Ascii characters representing an id number.
Eg: 0000000001

Length (size of the message): 10 bytes (10-19)
Right justified, zero filled, Ascii characters representing an id number.
Size is totally inclusive (n)
Eg: 0000000123

Device name: 64 bytes (20-83)
Left justified ASCII Characters representing a device name. Space filled.
Eg: HighVoltageSupply
CMD: 10 bytes (84-93)
Left justified ASCII Characters representing a device name. Space filled.
Eg: INVTHW____
Args: x bytes: 94-n
Left justified ASCII Characters representing any arguments to a command. Line feed terminated.
Eg: 10.123\n
ID of the message will always increment by one.
The response will contain the sender’s (i.e. whoever initiated the command) ID.

Instrument Controller Supported Commands
Unless otherwise stated, commands return an “OK” response when successful.  Failed commands return “FAIL” and unknown commands return “SYNTAX” in the response.

Instrument Controller Server Commands
Inventory Hardware
CMD:	INVTHW
Response:	INVTHW____[status string]\n
Description:	Multiple instances of the same command but with further information	filled in. Each response will have the device name field filled in with the	name of a device that is identified in the system XML configuration file. In	the arguments section will be a status string.

Shutdown System
CMD:	SHUTDOWN
Response:	None
Description:	Instrument controller will issue appropriate shutdown commands to all	devices and exit.

High Voltage Power Supply
Get Voltage and Current
CMD:	GETVI
Device Name:	HighVoltageSupply
Response:	OK________[voltage] [current]\n (check to see if it’s one	space or more)
Example:	GETV______\n
Description:	Will return the voltage (in volts) and current (in milliamps).

Set Voltage
CMD:	SETV [volts]
Device Name:	HighVoltageSupply
Response:	OK________\n (same thing here)
Example:	SETV______10.0\n
Description:	Will set the voltage to the specified amount (in volts).

Thermo Cycler (TEC)
Start Sequence
CMD:	STARTSEQ [number of cycles]
Device Name:	TECController
Response:	OK\n
Example:	STARTSEQ__40\n
Description:	Will begin the thermo cycler’s predefined steps for the given number	of cycles.
Stop Sequence
CMD:	STOPSEQ
Device Name:	TECController
Response:	OK\n
Example:	STOPSEQ___\n
Description:	Will abort the thermo cycler’s currently running sequence.

Read Sequence Data
CMD:	READSEQD
Device Name:	TECController
Response:	OK_[block temp]_[sample temp]_[current cycle]_[step]\n
Example:	READSEQD__\n
Description:	Will read data about the currently running sequence.  Infor returned will	be the block temperature in Celsius, the sample temperature in Celsius,	the current cycle number and the current step within the cycle.

OBIS Laser
Get Laser Power
CMD:	GETLPWR
Device Name:	OBISLaser
Response:	OK________[power in watts]\n
Example:	GETLPWR___\n
Description:	Will return the current output power (in watts).

Set Laser Power
CMD:	SETLPWR [watts]
Device Name:	OBISLaser
Response:	OK________\n
Example:	SETLPWR___0.01600\n
Description:	Will set the current output power (in watts).

Turn Laser ON or OFF
CMD:	SETLSTATE [ON|OFF]
Device Name:	OBISLaser
Response:	OK________\n
Example:	SETLSTATE_OFF\n
Description:	Will turn on or off the laser.

Laser’s Motor
Go Left
CMD:	MOVELEFT
Device Name:	LaserMotor
Response:	OK________\n
Example:	MOVELEFT\n
Description:	Laser Motor will go to the left

Go Right
CMD:	MOVERIGHT
Device Name:	LaserMotor
Response:	OK________\n
Example:	MOVERIGHT\n
Replace Capillary
CMD:	RETRACT
Device Name:	LaserMotor
Response:	OK________\n
Example:	RETRACT\n
Description:	Laser Motor retracts in order for user to put a new capillary.
Finished Replacing Capillary
CMD:	CAPREADY
Device Name:	LaserMotor
Response:	OK________\n
Example:	CAPREADY\n
Description:	Laser Motor moves back to it’s original position relative to capillary being	installed

Capillary Heater
Turn on Capillary Heater
CMD:	CAPHEATON
Device Name:	CapillaryHeater
Response:	OK________\n
Example:	CAPHEATON\n
Description: Turn on the Capillary Heater

Turn off Capillary Heater
CMD:	CAPHEATOFF
Device Name:	CapillaryHeater
Response:	OK________\n
Example:	CAPHEATOFF\n
Description:	Turn off the Capillary Heater

Get Temperature
CMD:	CAPGETT
Device Name:	CapillaryHeater
Response:	OK_____[temp]\n
Example:	CAPGETT\n
Description:	Get the temperature from the  Capillary Heater
Set Temperature
CMD:	CAPSETT [temp]
Device Name:	CapillaryHeater
Response:	OK_____\n
Example:	CAPSETT___10.0\n
Description:	Sets the temperature for the  Capillary Heater to heat up to.

Gel Pump
Push Gel Pump
CMD:	GELMV
Device Name:	GelPump
Response:	OK________\n
Example:	GELMV\n
Description:	Push Gel into capillary
Pull Gel Pump
CMD:	GELRET
Device Name:	GelPump
Response:	OK________\n
Example:	GELRET \n
Description:	Pull pump to refill Syringe with Gel

Gel Pump is filled and ready
CMD:	GELSTART
Device Name:	GelPump
Response:	OK________\n
Example:	GELSTART\n
Description: Pull pump to refill Syringe with Gel

Spectrometer
Set Exposure Time Milliseconds
CMD:	SPCSETEXP [timeMS]
Device Name:	Spectrometer
Response:	OK________\n
Example:	SPCSETEXP_250\n
Description:	Set the spectrometer exposure time in milliseconds.

Start Spectrometer continuous capture
CMD:	SPCSTARTC [filename] [time between reads MS] [capture duration MS]
Device Name:	Spectrometer
Response:	OK________\n
Example:	SPCSTARTC_spectra.csv 0 3600000\n
Description:	Start the spectrometer continuously capturing data to the given file. File will be created in the “data” folder. The filename should not have the “data” folder appended! The time between reads specifies how long to wait in between each capture during the specified duration. This can be used as a throttle to control the amount of data captured. The capture duration specifies for how long data should be captured. The example states that spectra will be captured to a file named spectra.csv that is located in the data folder with no delay between captrues and a duration of one hour.

Check if a continuous capture is running
CMD:	SPCISCRUN
Device Name:	Spectrometer
Response:	OK________[TRUE or FALSE]\n
Example:	SPCISCRUN\n
Description:	Checks if a continuous capture is running.

Fluidics Valve
Set Valve Postion
CMD:	FVALVEPOS [A|CLOSED|B]
Device Name:	FluidValve
Response:	OK________\n
Example:	FVALVEPOS_CLOSED\n
Description:	Control the position of the Fluidics valve. The valve has 3 postions	available: A, Closed and B.